External Interface for the OSTC3 Dive-Computer

Date: 121216
Last Update: 140501
Written by: Matthias Heinrichs (info@heinrichsweikamp.com)
Thanks to: Pascal Pelmont, Artur Wróblewski, Sven Knoch, Jef Driesen, JD Gascuel, Dirk Hohndel

Hardware
The USB interface is equipped with an USB Chip (FT232) from FTDI. You can get Windows/
Linux drivers from here: http://www.ftdichip.com/Drivers/VCP.htm (current Version is
2.08.24)
With Linux Kernel >2.6.9, these drivers are included in the Kernel.
In most cases, with Windows7 or Windows 8, these drivers will install automatically.
After installation a virtual COM port can be found in your windows hardware control panel.
You can change the COM port number there if required.

Software
The data rate is 115200 Baud, 8 Databits, 1 Stopbit, no parity, (115200, N, 8, 1)
16 and 24bit Values are transmitted in little Endian format (LOW:HIGH or (LOW:HIGH:UPPER)

Transfer protocol
Transfer is bi-directional and is initiated from the PC software by sending a command byte to the OSTC
- A command echo is send for every command
- Wait for answers/error bytes before starting new command!
- Keep COM port open until all transfers are done.
- After 120 seconds idle time the OSTC3 exits the COM mode – it's recommended to use a timeout lower then 120seconds and exit the COM mode from the PC application before the OSTC3 does (See Command 0xFF for details).
- Right button will abort COM mode when in idle loop.
- OSTC3 sends a 0x4D if it's ready for a new command.

Starting the communications with the OSTC3:
- [Open COM Port]
- Enter Download mode: 0xBB
- Answer from OSTC3:
- 0xBB (Command Echo)
- 0x4D -> OSTC3 will show "Download mode enabled"

Type	Command	Input	Output	Ready/Done.
Header	0x61	N	Data (65536)	0x4D
Clock	0x62	Data (6)	N	0x4D
Custom text	0x63	Data (60)	N	0x4D
Get Dive Profile	0x66	Data (1)	(Data n)	0x4D
Version/Identity	0x69	N	Data (64)	0x4D
Display Text	0x6E	Data (16)	N	0x4D
Reset all Settings	0x78	N	N	0x4D
Write Setting	0x77	See “Read/Write Settings”		0x4D
Read Setting	0x72	See “Read/Write Settings”		0x4D
Start Communication	0xBB	N	N	0x4D
Exit Communication	0xFF	N	N	N

command 0x61 („a“): download all 256 dive headers from the OSTC3
Output OSTC3:
0x61 (Command Echo)

The complete TOC (Table of Contents) memory block will be send, no matter how many dives are stored. One TOC entry is the header of one dive stored in the OSTC3's memory.
The Header does not contain the detailed profile for the dive.

The total of 256 headers (256 dives, max. in the OSTC3 Dive-Computer) will be send. No matter if all 256 dives are in the memory already used. Unused headers consists of 0xFF's only and must be ignored in the software.
Header size is 256byte/dive → A total of 65536Bytes will be transferred.
Download starts with dive #0 and ends with dive #255. Depending on how many dives made with the OSTC3 yet, this list is not sorted to the date and/or time of the dives. Sort by using date/time information from the header! Keep number (0-255) in your list, it's needed to get the dive profile for a particular dive. See command “f” for details.

Header details:
Byte Number	Length	Description	Data	Comment
0	2	Dive Header start	0xFA, 0xFA	For internal use
2	3	Internal pointer to the begin of the profile data	24bit Address	For internal use
5	3	Internal pointer to the end of the profile data	24bit Address	For internal use
8	1	Logbook-Profile version	8bit Version	Currently only 0x23
9	3	Length of profile data for this dive	24bit Length	Useful for parsing profile
12	3	Date of Dive (YYMMDD)	3Byte Date	Year is 20xx
15	2	Time at end of Dive (HHMM)	2Byte Time
17	2	Max. Depth in mbar (Includes salinity offset!)	16bit Depth	See Note 1
19	3	Divetime (MMMMSS)	24bit Divetime	2Bytes Min, 1Byte Sec.
22	2	Min. Water temperature in 0.1°C	16bit Temp	Signed 16Bit
24	2	Surface air pressure before dive in mbar	16bit Pressure	30min before dive
26	2	Desaturation time after the dive in minutes	16bit Minutes
28	4	Gas/Dil 1 (%O2, %He, Change Depth in m, Type)	4Byte Gas data	See Note 2,3
32	4	Gas/Dil 2 (%O2, %He, Change Depth in m, Type)	4Byte Gas data	See Note 2,3
36	4	Gas/Dil 3 (%O2, %He, Change Depth in m, Type)	4Byte Gas data	See Note 2,3
40	4	Gas/Dil 4 (%O2, %He, Change Depth in m, Type)	4Byte Gas data	See Note 2,3
44	4	Gas/Dil 5 (%O2, %He, Change Depth in m, Type)	4Byte Gas data	See Note 2,3
48	2	Firmware Version xx.yy	2Byte Version
50	2	Battery voltage in mV after the dive	16bit Voltage
52	1	Sampling rate in Seconds	1Byte	Also stored in profile section
53	2	CNS at beginning of dive in %	16bit CNS
55	1	GF at beginning of dive in %	1Byte GF
56	1	GF at end of dive in %	1Byte GF
57	2	Configured logbook offset in the OSTC3	16Bit Offset
59	1	Unused	1Byte	Ignore
60	2	Setpoint 1 (Setpoint in cbar, Change depth in m)	2Byte SP Data	1 cbar = 0.01bar
62	2	Setpoint 2 (Setpoint in cbar, Change depth in m)	2Byte SP Data
64	2	Setpoint 3 (Setpoint in cbar, Change depth in m)	2Byte SP Data
66	2	Setpoint 4 (Setpoint in cbar, Change depth in m)	2Byte SP Data
68	2	Setpoint 5 (Setpoint in cbar, Change depth in m)	2Byte SP Data
70	1	Salinity Setting	1Byte Salinity	See Note 5
71	2	Max. CNS in %	16bit CNS
73	2	Average depth in mbar	16bit Depth	See Note 1
75	2	Total dive time in seconds (incl. shallow parts)	16bit Seconds	See Note 6
77	1	If GF deco model was used : GF_lo [%] / else: Saturation multiplier [%]	1Byte Deco model info
78	1	If GF deco model was used : GF_hi [%] / else: Desaturation multiplier [%]	1Byte Deco model info
79	1	Deco model used for this dive	1Byte Model	0= ZH-L16, 1= ZH-L16-GF
80	2	Dive number (Internal total dive counter)	16bit Number	See Note 7
82	1	Dive mode used for this dive	1Byte Mode	See Note 8
83	16	N2 compartment desaturation time in min.	16Bytes	See Note 9
99	64	N2 compartments at end of dive in mbar	64Bytes	16*4Byte (Float), See Note 9
163	16	He compartment desaturation time in min.	16Bytes	See Note 9
179	64	He compartments at end of dive in mbar	64Bytes	16*4Byte (Float), See Note 9
243	1	Last decompression stop in m	1Byte Depth	User-Set value
244	1	Assumed distance to shown stop in 10cm	1Byte	Safety-margin parameter
245	2	hwHUD battery value at end of dive in mV	16bit Voltage	Zero when no HUD present
247	1	hwHUD last status	1Byte Status	See Note 10
248	6	Battery gauge registers at end of dive in 1*10-9As	48bit Battery	Only for 3,6V Battery
254	2	Dive Header end	0xFB, 0xFB

Note 1: Depth in meters = Depth in mbar*g0/1000 (g0=9,80665m/s2)
Note 2: “First Gas” always has change depth = 0m, List contains OC Gases for OC and Diluents for CC
Note 3: Gas Types: 0=Disabled, 1=First, 2= Travel, 3= Deco, Diluent Type: 0=Disabled, 1=First, 2=Normal
Note 4: bit0-3: Divisor, 0=never, 1=every Sample, 2=every 2nd Sample, ...bit4-7: Information size (0-15)
Note 5: 0-5 [%]
Note 6: very shallow (<1m) parts of the dive are not counted in the divetime (Byte 19-21)
Note 7: This number is increased by one with every dive made and the number stored here is the number of total dives at the time this dive was made. It's not related to the TOC entry.
Note 8: 0= OC, 1= CC, 2= Gauge, 3= Apnea
Note 9: Compartment assignment according to Bühlmann, Albert: Tauchmedizin; 4. Auflage [2002], values after dive
Note 10: bit0=1:HUD connection ok, bit1=1:HUD is calibrated , bit2=1:HUD Battery is low (<3000mV), bit3=1:Sensor 1 active, bit4=1:Sensor 2 active, bit5=1:Sensor 3 active, bit6=1:HUD calibrated through external interface (DR5/DRx), bit7=1:HUD received (valid) data through RS232 input from DR5/DRx

Note: OSTC3 sends a 0x4D after transferring all headers successfully.
Profile Details

The profile for each dive has a variable length.

First, a small header with the information divisors is send.
Small Header
Byte Number	Length	Description	Data	Comment
0	3	Length of profile data for this dive	24bit Length	Useful for parsing profile
3	1	Sampling rate in Seconds	1Byte	Also stored in header section
4	1	Number of Divisors in this dive (n=7)	1Byte	Each divisor has 3 Bytes, See Note 1
5 (1*3+2)	3	Sampling Divisor Temperature	3Byte	See Note 2
8 (2*3+2)	3	Sampling Divisor Deco/NDL Status	3Byte	See Note 2
11 (3*3+2)	3	Sampling Divisor Gradient Factor	3Byte	See Note 2
14 (4*3+2)	3	Sampling Divisor ppO2 Sensors	3Byte	See Note 2
17 (5*3+2)	3	Sampling Divisor Decoplan	3Byte	See Note 2
20 (6*3+2)	3	Sampling Divisor CNS profile	3Byte	See Note 2
23 (n*3+2)	3	Sampling Divisor Tank data	3Byte	See Note 2
(n+1*3+2)		Future Profile versions may have more divisors

Note 1: Inspect each divisor (0=no Divisors, 1=One Divisor, etc.)

Note 2: Each divisor: Byte0= Type, Byte1=Length, Byte2=Divisor
Type=0: Temperature
Type=1: Deco/NDL Status
Type=2: Gradient Factor
Type=3: ppO2 Sensors
Type=4: Decoplan
Type=5: CNS
Type=6: Tank Data

Now, the actual dive profile is send.
Byte (n*3+5) to Byte #“Length of profile data for this dive”:

Depth in mbar (LOW:HIGH, not including salinity offset), ProfileFlagByte, extended Information
...
2 Byte:
0xFD 0xFD (Profile and dive end)

ProfileFlagByte:
bit0-6: amount of additional bytes after the ProfileFlagByte (0-127)
bit7: EventByte Flag (At least one Event Byte is available)

EventByte:
bit0-3: Alarm type
0=No Alarm
1=unused
2=Deco Stop missed
3=unused
4=ppO Low Warning
5=ppO High Warning
6=Manual Marker
7=Low Battery
Bit4: Manual Gas Set („GasX”)
Bit5: Gas Change (Gas1-5 or Dil 6-10)
Bit6: SetPoint Change (cbar)
Bit7: Another Eventbyte is available

EventByte 2 (Only included if Bit7 of Eventbyte 1 = 1)
Bit0: Switched to bailout
Bit1-6: unused (yet)
Bit7:Another Eventbyte is available...

Events and Extended Information will be stored strictly in the following order (future extensions will append at the end of the list for compatibility reasons!):
Events:
- EventByte (s)
- ManualGasSet: new values for Gas 6 (2 bytes: byte1=O2%, byte2=He%)
- Gas Change (1 byte: current Gas#)
- SetPoint Change (1Byte: new SetPoint in cbar)
- Bailout: new values for bailout gas (2 bytes: byte1=O2%, byte2=He%)

Extended Information:
Type	Length	Data	Standard Divisor	Comment
Temperature	2	16bit short signed in 0.1°C	6
Deco/NDL Status	2	Byte1:=0 if within NDL, if in deco mode: depth of 1st stop / Byte2: NDL in minutes or length of first stop in minutes)	6
GF Value	1	1Byte GF in %	12
PpO2-Sensors	9	Per Sensor three bytes. / Byte0: Sensor ppO2 (in 0.01bar) / Byte1-2: 16bit Sensor Voltage in 0.1mV steps	2	Divisor=0 in OC dives
Deco Plan	15	Last 15 stops in minutes	12	Make sure to take care of header Byte243, which can be 3, 4, 5 or 6m
CNS value	2	16bit CNS value in %	12
Tank Data	0	Not used yet	0

For calculation of the available information the following rules apply:
- First sample# is 1
- if divisor == 0 the information will never appear
- dataAvailable = (sample# modulo divisor) == 0

Sampling rate can vary between dives! Always use sample rate from Dive Header.

Profile memory is a „first-in first-out” memory, so older dives will be overwritten when there is no free space left. Profile for dives, where no profile information is available, will be two bytes: 0xFD 0xFD

Examples:
Assume the following configuration:
Divisor Temperature=3
Divisor Deco/NDL Status=1
Divisor Gradient Factor=0
Divisor ppO2 Sensors=0
Divisor Decoplan=0
Divisor CNS=0
Divisor Tank Data=0

Sample 1:
0x64 0x00 0x02 0x00 0xA0
Depth = 1m
ProfileFlagByte = 2 bytes will follow, No Event
Deco/NDL = NDL 160 minutes

Sample 2:
0xC8 0x00 0x02 0x00 0xA0
Depth = 2m
ProfileFlagByte = 2 bytes will follow, No Event
Deco/NDL = NDL 160 minutes

Sample 3:
0x2C 0x01 0x04 0x32 0x00 0x00 0xA0
Depth = 3m
ProfileFlagByte = 4 bytes will follow, No Event
Temperature = 5°C
Deco/NDL = NDL 160 minutes

Sample 4:
0x5E 0x01 0x84 0x21 0x02 0x00 0x9f
Depth = 3.5m
ProfileFlagByte = 4 bytes will follow, Event available (bit7=1)
EventByte = Alarm SLOW, Stored Gas Change
GasChange = Gas#2
Deco/NDL = NDL 159 minutes
...

Other Commands
command 0x62 („b“): Sync clock and date with PC
Output OSTC3: 0x62 (command echo)
Now, 6 bytes must be send to the OSTC3
Byte1: Hours
Byte2: Minutes
Byte3: Seconds
Byte4: Month
Byte5: Day
Byte6: Year (20xx)
All bytes will be checked for plausibility and the clock is set at once. After a timeout of about 30ms, the transmission will be aborted
Note: OSTC3 sends a 0x4D after setting the clock successfully

command 0x63 („c“): Send the custom text (up to 60 bytes)
Output OSTC3: 0x63 (command echo)
Now, up to 60 bytes must be send to the OSTC3
Byte 1-60: Custom text string shown when in surface mode (NULL terminated, last byte =0x00)
After a timeout of about 30ms, the transmission will be aborted and the text saved.
Note: OSTC3 sends a 0x4D after the text is saved.

command 0x66 („f“): Download one dive (incl. Profile data)
Output OSTC3: 0x66 (command echo)
Now, 1 byte must be send to the OSTC3
Byte1: Number of dive to download (0-255)
This number is the internal TOC number in which the OSTC3 has stored this dive.

Now, the OSTC3 sends the header (256bytes) + the profile (Ending with 0xFD 0xFD).
Note: Header bytes 8-10 contain the profile length in bytes including final 0xFD 0xFD
Depending on the dive length and profile details you can expect about 3Bytes/second divetime (e.g. 21,6kByte for a one-hour dive).

command 0x69 („i“): Download OSTC3 identity
Output OSTC3: 0x69 (command echo)
Output OSTC3: serial number (low:high)
Output OSTC3: firmware version (x.y)
Output OSTC3: Custom text (60 chars) Note: Custom Text is NULL terminated, but here all 60bytes are send.
Output OSTC3: Final 0x4D

command 0x6E („n“): Display up to 16 characters in the COM mode
Output OSTC3: 0x6E (command echo)
Now, up to 16 bytes must be send to the OSTC3
Byte 1-16: String to show in COM mode
Note: This is useful for debugging and showing some status (e.g.: “Reading Dives”) before actually downloading data
(Which may take up to 2minutes for a full memory download).
Text can be updated as often as required whenever the COM mode is in the idle state waiting for commands. As all other commands, this will also reset the timeout.
Note: OSTC3 sends a 0x4D after displaying the new text

command 0xFF: Exit COM mode
Output OSTC3: 0xFF (command echo)
The OSTC3 displays “Exited” for two seconds and powers-down the COM chip.

Command 0x77 and 0x72: Read/Write Settings:

1. Byte (Read/Write)
Write: 0x77 („w“)
 +2...5 Bytes Data
OSTC3 will send 0x4D when all bytes have been recieved
Read: 0x72 („r“)
 +1 Byte Data
OSTC3 will send 1...4 Bytes information + 0x4D
Output OSTC3: 1 Byte (command echo)

2. Byte	3. Byte	4. Byte	5. Byte	6. Byte
Gas1: 0x10	%O2	%He	Type (Note 1)	Change Depth [m]
Gas2: 0x11	%O2	%He	Type (Note 1)	Change Depth [m]
Gas3: 0x12	%O2	%He	Type (Note 1)	Change Depth [m]
Gas4: 0x13	%O2	%He	Type (Note 1)	Change Depth [m]
Gas5: 0x14	%O2	%He	Type (Note 1)	Change Depth [m]
Dil1: 0x15	%O2	%He	Type (Note 2)	Change Depth [m]
Dil2: 0x16	%O2	%He	Type (Note 2)	Change Depth [m]
Dil3: 0x17	%O2	%He	Type (Note 2)	Change Depth [m]
Dil4: 0x18	%O2	%He	Type (Note 2)	Change Depth [m]
Dil5: 0x19	%O2	%He	Type (Note 2)	Change Depth [m]
SP1: 0x1A	Setpoint [cbar]	Change Depth [m]
SP2: 0x1B	Setpoint [cbar]	Change Depth [m]
SP3: 0x1C	Setpoint [cbar]	Change Depth [m]
SP4: 0x1D	Setpoint [cbar]	Change Depth [m]
SP5: 0x1E	Setpoint [cbar]	Change Depth [m]
CCR Mode: 0x1F	0=Fixed SP, 1=Sensor
Dive Mode: 0x20	0=OC, 1=CC, 2=Gauge, 3=Apnea
Deco type: 0x21	0=ZH-L16, 1=ZH-L16-GF
ppO2 max: 0x22	ppO2 max [cbar]
ppO2 min: 0x23	ppO2 min [cbar]
Future TTS: 0x24	FTTS [min]
GF_low: 0x25	GF_low [%]
GF_high: 0x26	GF_high [%]
aGF_low: 0x27	aGF_low [%]
aGF_high: 0x28	aGF_high [%]
aGF selectable: 0x29	0=aGF disabled, 1=aGF can be selected underwater
Saturation: 0x2A	Saturation [%]
Desaturation: 0x2B	Desaturation [%]
Last Deco: 0x2C	Last Stop [m]
Brightness: 0x2D	0=Eco, 1=Medium, 2=Full
Units: 0x2E	0=m/°C, 1=ft/°F
Sampling rate: 0x2F	0=2s, 1=10s
Salinity: 0x30	0-5 [%]
Dive Mode Colour: 0x31	0=Standard, 1=Red, 2=Green, 3=Blue
Language: 0x32	0=EN, 1=DE, 2=FR, 3=ES
Date Format: 0x33	0=MMDDYY, 1=DDMMYY, 2=YYMMDD
Compass Gain: 0x34	(Note 3)
Pressure Sensor Offset: 0x35	In mbar (Note 4)

Note 1: 0=Disabled, 1=First, 2=Travel, 3=Deco
Note 2: 0=Disabled, 1=First, 2=Normal
Note 3: 0=230LSB/Gauss, 1=330LSB/Gauss, 2=390LSB/Gauss, 3=440LSB/Gauss, 4=660LSB/Gauss, 5=820LSB/Gauss, 6=1090LSB/Gauss, 7=1370LSB/Gauss (See HMC5883L Datasheet, Page 13, for details)
Note 4: signed int8 (two's complement), values between +/-20 are valid

Note: OSTC3 sends a 0x4D after read or write of one setting.
Note: Use command 0x6E („n“): Display up to 16 characters in the COM mode to display some feedback on the OSTC3 screen when updating settings.

Example 1: Set Gas 3 to TX18/40 (First Gas)
Command: 0x77 (Write)
Output OSTC3: 0x77 (command echo)
Command: 0x12, d'17', d'40', 0x01, 0x00 (Gas3, O2=17, He=40, Type=1, Change Depth=0)
Output OSTC3: 0x4D (Command done)

Example 2: Set Gas 5 to TX50/10 (Deco Gas, 21m)
Command: 0x77 (Write)
Output OSTC3: 0x77 (command echo)
Command: 0x14, d'50', d'10', 0x03, d'21' (Gas5, O2=50, He=10, Type=3, Change Depth=21)
Output OSTC3: 0x4D (Command done)

Example 3: Set OSTC3 to CCR Mode with Fixed Setpoints
Command: 0x77 (Write)
Output OSTC3: 0x77 (command echo)
Command: 0x20, 0x01 (Dive mode=1)
Output OSTC3: 0x4D (Command done)
Command: 0x77 (Write)
Output OSTC3: 0x77 (command echo)
Command: 0x1F, 0x00 (CCR Mode=0)
Output OSTC3: 0x4D (Command done)

As soon as command 1 is done (0x4D from OSTC3) the next command can be started.

command 0x78 („x“): Reset all settings
Output OSTC3: 0x78 (command echo)
Output OSTC3: Final 0x4D