External Interface for the Open-Source Dive-Computer (OSTC, OSTC Mk.2/2N)

Date: 01/12/08
Last Update: 04/28/12 (Valid for 2.21stable and later)
Written by: Matthias Heinrichs (info@heinrichsweikamp.com) and Pascal Pellmont

Hardware
The USB interface is equipped with an USB Chip (FT232) from FTDI. You can get Windows/
Linux drivers from here: http://www.ftdichip.com/Drivers/VCP.htm (current Version is
2.08.24)
With Linux Kernel >2.6.9, these drivers are included in the Kernel.
With Windows7 these drivers will install automatically.
After installation a virtual COM port can be found in your windows hardware control panel.
You can change the COM port number there if required.

Software
The data rate is 115200 Baud, 8 Databits, 1 Stopbit, no parity, (115200, N, 8, 1)
16bit Values are transmitted in little Endian format (LOW:HIGH)

Transfer protocol
Transfer is bi-directional and is initiated from the PC software by sending a command byte to
the OSTC

command 0x61 („a“): download status and dive profiles from OSTC
Output OSTC:
Byte1-5:
0xAA 0xAA 0xAA 0xAA 0xAA (Preamble)
Byte6:
0x55 (Preamble)
Byte7-262:
Internal EEPROM BANK 0 (First 256Bytes) See Command „g”
Byte263-264:
current battery voltage in mV
Byte265-266:
2 Byte current firmware version , 1st Byte.2nd Byte (e.g. „1.90“)

With firmware version 1.90 or lower (32kB Logbook):
Byte267-33034 (32768 Byte):
logbook with profile data
With firmware version 1.91 or higher (64kB Logbook):
Byte267-65802 (65536 Byte):
logbook with profile data

The complete profile memory will be send, no matter how many dives are stored.
One dive consists of a header and the profile block. There may be two different headers in the data stream. Byte 2 (Logbook-Profile version) may be 0x20 or 0x21:
0x20 → The header is send with a length of 47 bytes
0x21 → The header is send with a length of 57 bytes (Additional 10 bytes at the end of the header)
Requirements for the parser to be compatible with all firmware versions: / Check Byte 265-266 for the used firmware: / Firmware 1.90 or lower → Read additional 32768Byte logbook profile data / Firmware 1.91 or higher → Read additional 65536Byte logbook profile data / For each dive found in the logbook profile data: / Check Byte2 (Logbook-Profile version) for 0x20 or 0x21 / 0x20 → Dive has a 47Byte header / 0x21 → Dive has a 57Byte header with some extra information not included in the 0x20 version

„0x20” Details:
Byte 0-1:
0xFA 0xFA (start Header)
Byte 2:
Logbook-Profile version (0x20 for this 47 Byte version)
Byte 3:
Date of dive, month
Byte 4:
Date of dive, day
Byte 5:
Date of dive, year (20xx)
Byte 6:
Time at end of dive, hours
Byte 7:
Time at end of dive, minutes
Byte 8-9:
max. depth in mbar (Includes salinity offset!)
Byte 10-11:
divetime (minutes)
Byte 12:
divetime (seconds)
Byte 13-14:
min. water temperature in 0.1 °C
Byte 15-16:
Surface air pressure before dive
Byte 17-18:
Desaturation time in minutes after the dive
Byte 19-20:
Gas 1 current (%O2, %He)
Byte 21-22:
Gas 2 current (%O2, %He)
Byte 23-24:
Gas 3 current (%O2, %He)
Byte 25-26:
Gas 4 current (%O2, %He)
Byte 27-28:
Gas 5 current (%O2, %He)
Byte 29-30:
Gas 6 current (%O2, %He) (Gas 6 is a manual gas that can be configured during the dive),
value at end of dive.
Byte 31:
Selected Gas# at beginning of dive
Byte 32-33:
2 Byte firmware version , 1st Byte.2nd Byte (e.g. „1.90“)
Byte 34-35:
2 Byte battery voltage in mV after the dive
Byte 36:
Sampling Rate in Seconds
Byte 37:
Sampling Divisor Temperature
Bit 0-3: Divisor, 0=never, 1=every Sample, 2=every 2nd Sample, ...
Bit 4-7: Information size (currently 2)
Byte 38:
Sampling Divisor Deco/NDL Status
Bit 0-3: Divisor, 0=never, 1=every Sample, 2=every 2nd Sample, ...
Bit 4-7: Information size (currently 2)
Byte 39:
Sampling Divisor Gradient Factor
Bit 0-3: Divisor, 0=never, 1=every Sample, 2=every 2nd Sample, ...
Bit 4-7: Information size (currently 1)
Byte 40:
Sampling Divisor ppO2 Sensors
Bit 0-3: Divisor, 0=never, 1=every Sample, 2=every 2nd Sample, ...
Bit 4-7: Information size (currently 3)
Note: The OSTC Mk.2 does not support ppO2 Sensors.
Byte 41:
Sampling Divisor “Deco Debug”
Bit 0-3: Divisor, 0=never, 1=every Sample, 2=every 2nd Sample, ...
Bit 4-7: Information size (currently 9)
Byte 42:
Sampling Divisor CNS profile
Bit 0-3: Divisor, 0=never, 1=every Sample, 2=every 2nd Sample, ...
Bit 4-7: Information size (currently 1)
Byte 43:
Salinity Setting
 d‘100‘ = 1.00kg/l
 d‘101‘ = 1.01kg/l
 d‘102‘ = 1.02kg/l
 d‘103‘ = 1.03kg/l
 d‘104‘ = 1.04kg/l
Byte 44:
Max. CNS in % (For dives made with Firmware 1.84 and later)
Byte 45-46:
0xFB 0xFB (end of header)

„0x21” Details:
Byte 0-1:
0xFA 0xFA (start Header)
Byte 2:
Logbook-Profile version (0x21 for this 57 Byte version)
Byte 3-44:
Identical with „0x20” version of header description
Byte 45-46:
Average depth in mbar
Byte 47-48:
Total dive time in seconds (Regardless of CF01 - shallow parts of the dive are not counted in the divetime (Byte 10-12)
Byte 49:
If GF deco model was used for this dive: GF_lo in percent (CF32)
If non-GF (Standard Bühlmann) was used: Saturation multiplier in percent (CF11)
Byte 50:
If GF deco model was used for this dive: GF_hi in percent (CF33)
If non-GF (Standard Bühlmann) was used: Desaturation multiplier in percent (CF12)
Byte 51: Deco type used for this dive: (0=ZH-L16 OC, 1=Gauge, 2=ZH-L16 CC, 3=Apnoe 4=L16-GF OC, 5=L16-GF CC, 6=PSCR-GF )
Byte 52-54:
Not used/Reserved =0x00
Byte 55-56:
0xFB 0xFB (end of header)

After the header, the profile is send. It has a variable length and is identical for „0x20” or „0x21” Headers:

Depth in mbar (LOW:HIGH, not including salinity offset), ProfileFlagByte, extended Information
...
2 Byte:
0xFD 0xFD (Profile and dive end)

ProfileFlagByte:
Bit0-6: amount of additional bytes after the ProfileFlagByte (0-127)
Bit7: EventByte Flag (Event Byte is available)

EventByte:
Bit0-3: Alarm type
0=No Alarm
1=SLOW
2=Deco Stop missed/GF violated
3=Deep Stop missed
4=ppO Low Warning
5=ppO High Warning
6=Manual Marker
7=Low Battery
...
Bit4: Manual Gas Set („Gas 6”)
Bit5: Gas Change (Gas1-5)
Bit6: SetPoint Change (cbar)
Bit7: unused

Extended Information will be stored strictly in the following order (future extensions will append at the end of the list for compatibility reasons!):
- EventByte (1 byte)
- ManualGasSet: new Values for Gas 6 (2 bytes: byte1=O2%, byte2=He%)
- Gas Change (1 byte: current Gas#)
- Temperature in 0.1°C (2 bytes: LOW:HIGH)
- Deco/NDL Status (2 bytes: byte1: =0 if within no deco limit, if in deco mode: depth of 1st stop; byte2: NDL in minutes or length of first stop in minutes)
- GF value (1Byte: GF value in %)
- ppO2 Sensor Values (according to header information, t.b.d)
- Deco Debug (according to header information)
- CNS value (1Byte: CNS value in %)
- SetPoint Change (1Byte: new SetPoint in cbar)

For calculation of the available information the following rules apply:
- First sample# is 1
- if divisor == 0 the information will never appear
- dataAvailable = (sample# modulo divisor) == 0

Sampling rate can vary between dives! Always use sample rate from Dive Header.

Profile memory is a „first-in first-out” memory, so older dives will be overwritten when there is no free space left. This can result in „half” or broken dives in the memory. A complete dive always has a complete Header (0xFA 0xFA to 0xFB 0xFB) and a complete profile block (ending with 0xFD 0xFD) The latest dive has an additional 0xFE after the profile block. This byte is required to find the latest dive after a firmware update, etc.

Examples:
Assume the following configuration:
Divisor Temperature=3
Divisor Deco/NDL=1
Divisor GF=0
Divisor ppO2=0
Divisor Deco debug=0
Divisor CNS=0.

Sample 1:
0x64 0x00 0x02 0x00 0xA0
Depth = 1m
ProfileFlagByte = 2 bytes will follow, No Event
Deco/NDL = NDL 160 minutes

Sample 2:
0xC8 0x00 0x02 0x00 0xA0
Depth = 2m
ProfileFlagByte = 2 bytes will follow, No Event
Deco/NDL = NDL 160 minutes

Sample 3:
0x2C 0x01 0x04 0x32 0x00 0x00 0xA0
Depth = 3m
ProfileFlagByte = 4 bytes will follow, No Event
Temperature = 5°C
Deco/NDL = NDL 160 minutes

Sample 4:
0x5E 0x01 0x84 0x21 0x02 0x00 0x9f
Depth = 3.5m
ProfileFlagByte = 4 bytes will follow, Event available (Bit7=1)
EventByte = Alarm SLOW, Gas Change
GasChange = Gas#2
Deco/NDL = NDL 159 minutes
...

command 0x62 („b“): Sync clock and date with PC
Output OSTC: 0x62
Now, 6 bytes must be send to the OSTC
Byte1: Hours
Byte2: Minutes
Byte3: Seconds
Byte4: Month
Byte5: Day
Byte6: Year (20xx)
All bytes will be checked for plausibility and the clock is set at once. After a timeout of about 30ms, the transmission will be aborted

command 0x63 („c“): Starts simulator mode
This command sets the OSTC into simulator mode. Every byte following this command will set ambient_pressure to [byte]*100mbar and the OSTC will operate as if the unit is actually in use underwater e.g.: 0x1E (dec 30) simulates a pressure of 30*100mbar=3bar=20m depth
Simulator mode will be aborted if the device lefts the divemode after the simulated dive. (ascend to 0m by sending 0x00 and wait)

command 0x64 („d“): Write access to internal EEPROM BANK 0
This command allows write access to the internal EEPROM from address 0x004 to 0x0FF (Serial number and total dives can not be written): The first four bytes of EEPROM Bank0 are not available for writing.
Output OSTC: 0x64
Now, every incoming byte will be written to the internal EEPROM starting with address 0x004. Each byte will be echoed if the write was completed.
Important: Saving one byte takes about 4ms, so the next byte must be send AFTER the echo of the current byte was received from the software.
A total of 251Byte can be send. After a timeout of about 30ms, the current byte transmission will be aborted and 0xFF is send instead of the received value

Warning: EEPROM Bank0 and Bank1 include important settings, e.g. the battery history. Make sure to re-write the original data when changing settings!

command 0x65 („e“): Send MD2 Hash fingerprint to PC
Byte1-2:
2 Byte firmware version , 1st Byte.2nd Byte (e.g. „1.90“)
Byte3-18:
16 Byte MD2 hash fingerprint (0x00 for all bytes if hash was not built yet)

command 0x66 („f“): Unused/Reserved

command 0x67 („g“): Send internal EEPROM BANK 0 only
This command can be used to read out the gas setup and custom functions without reading the profile data (faster download)
Byte1-2:
Serial number(Not writable!)
Byte3-4:
Total dives(Not writable!)
Byte5-6:
Gas 1 default (%O2=21, %He=0)
Byte7-8:
Gas 1 current (%O2, %He)
Byte9-10:
Gas 2 default (%O2=21, %He=0)
Byte11-12:
Gas 2 current (%O2, %He)
Byte13-14:
Gas 3 default (%O2=21, %He=0)
Byte15-16:
Gas 3 current(%O2, %He)
Byte17-18:
Gas 4 default (%O2=21, %He=0)
Byte19-20:
Gas 4 current (%O2, %He)
Byte21-22:
Gas 5 default (%O2=21, %He=0)
Byte23-24:
Gas 5 current (%O2, %He)
Byte25-26:
Gas 6 current (%O2, %He)
Byte27:
Salinity Setting
 d‘100‘ = 1.00kg/l
 d‘101‘ = 1.01kg/l
 d‘102‘ = 1.02kg/l
 d‘103‘ = 1.03kg/l
 d‘104‘ = 1.04kg/l
(Other values then these four are ignored due to the CE depth accuracy requirements)
Byte28:
Active Gas Flag Register
 Bit0=Gas1
 Bit1=Gas2
 Bit2=Gas3
 Bit3=Gas4
 Bit4=Gas5
 Bit5=not used/reserved
 Bit6: not used/reserved
 Bit7: not used/reserved
Byte29:
Gas1: Change depth in m
Byte30:
Gas2: Change depth in m
Byte31:
Gas3: Change depth in m
Byte32:
Gas4: Change depth in m
Byte33:
Gas5: Change depth in m
Byte34:
First Gas number (1-5)
Byte35:
DecoType (0=ZH-L16 OC, 1=Gauge, 2=ZH-L16 CC, 3=Apnoe 4=L16-GF OC, 5=L16-GF CC, 6=PSCR-GF)
Byte36:
Use O2 Sensor Module in CC Modes (0= OFF, 1= ON) (Only available in old OSTC1 - unused for OSTC Mk.2/2N)
Byte37:
SetPoint1: ppO2 (x 0.01bar, e.g. 155 = 1.55bar ppO2)
Byte38:
SetPoint2: ppO2 (x 0.01bar, e.g. 155 = 1.55bar ppO2)
Byte39:
SetPoint3: ppO2 (x 0.01bar, e.g. 155 = 1.55bar ppO2)
Byte40:
Debugmode (0= OFF, 1= ON). Using the debug mode is not recommended for end-users.
Byte41-42:
Lowest Battery voltage seen (in mV)
Byte43:
Lowest Battery voltage seen at (Month)
Byte44:
Lowest Battery voltage seen at (Day)
Byte45:
Lowest Battery voltage seen at (Year)
Byte46-47:
Lowest Battery voltage seen at (Temperature in 0.1 °C)
Byte48:
Last complete charge at (Month)
Byte49:
Last complete charge at (Day)
Byte50:
Last complete charge at (Year)
Byte51-52:
Total charge cycles
Byte53-54:
Total complete charge cycles
Byte55-56:
Temperature Extrema minimum (Temperature in 0.1 °C)
Byte57:
Temperature Extrema minimum at (Month)
Byte58:
Temperature Extrema minimum at (Day)
Byte59:
Temperature Extrema minimum at (Year)
Byte60-61:
Temperature Extrema maximum (Temperature in 0.1 °C)
Byte62:
Temperature Extrema maximum at (Month)
Byte63:
Temperature Extrema maximum at (Day)
Byte64:
Temperature Extrema maximum at (Year)
Byte65:
Custom Text active (=1), Custom Text Disabled (<>1)
Byte66-90:
(25Bytes): Custom Text for Surfacemode (Real text must end with "}")
Example: OSTC Dive Computer} (19 Characters incl. "}") Bytes 85-90 will be ignored.
Byte91:
Dim OLED in Divemode (>0), Normal mode (=0)
Byte92:
Date format for all outputs:
=0: MM/DD/YY
=1: DD/MM/YY
=2: YY/MM/DD
Byte93:
Total number of CF used in installed firmware
Byte94-128
not used/reserved
Byte129-256:
32 custom Functions (CF0-CF31)
The custom functions are stored in the internal EEPROM after 0x080
Any custom function occupies 4 byte:
2 byte (low:high) store the default value, reset from menu "reset" (Will be overwritten by hard-coded defaults if reset in the unit!)
2 byte (low:high) store the actual value

command 0x68 („h“): Reset DecoData
Output OSTC: 0x68 (command echo)

command 0x69 („i“): Write access to internal EEPROM BANK 1
This command allows write access to the internal EEPROM from address 0x104 to 0x1FF.
The first four bytes of EEPROM Bank1 are not available for writing.
Output OSTC: 0x69
Now, every incoming byte will be written to the internal EEPROM starting with address 0x104. Each byte will be echoed if the write was completed.
Important: Saving one byte takes about 4ms, so the next byte must be send AFTER the echo of the current byte was received from the software.
A total of 251Byte can be send. After a timeout of about 30ms, the current byte transmission will be aborted and 0xFF is send instead of the received value

Warning: EEPROM Bank0 and Bank1 include important settings, e.g. the battery history. Make sure to re-write the original data when changing settings!

command 0x6A („j“): Send internal EEPROM BANK 1
This command is required to read out the BANK1 of the EEPROM
Byte1:
Logbook version indicator (Not writable!)
Byte2-3:
Last Firmware installed, 1st Byte.2nd Byte (e.g. „1.90“) (Not writable!)
Byte4:
OLED brightness (=0: Eco, =1 High) (Not writable!)
Byte5-128
not used/reserved
Byte129-256:
32 custom Functions (CF 32-63)
The custom functions are stored in the internal EEPROM after 0x180
Any custom function occupies 4 byte:
2 byte (low:high) store the default value, reset from menu "reset" (Will be overwritten by hard-coded defaults if reset in the unit!)
2 byte (low:high) store the actual value

command 0x6B („k“): Save tissue data and restore after simulation
Output OSTC: 0x6B (command echo)
Tissue data is saved and restored after the simulation. Next simulated dive after this command will not affect tissue data. Useful if remaining saturation data should not affected from a PC-simulated dive.

command 0x6C („l“): Dump screen shot
Output OSTC: 0x6C (command echo)
A screen shot is send via the UART interface in a compressed format. Contact for details.

command 0x6D („m“): Send internal EEPROM BANK 2
This command is required to read out the BANK2 of the EEPROM
Byte1-4:
not used/reserved (Not writable!)
Byte5-128:
not used/reserved
Byte129-256:
32 custom Functions (CF 64-95)
The custom functions are stored in the internal EEPROM after 0x280
Any custom function occupies 4 byte:
2 byte (low:high) store the default value, reset from menu "reset" (Will be overwritten by hard-coded defaults if reset in the unit!)
2 byte (low:high) store the actual value

command 0x6E („n“): Write access to internal EEPROM BANK 2
This command allows write access to the internal EEPROM from address 0x204 to 0x2FF.
The first four bytes of EEPROM Bank2 are not available for writing.
Output OSTC: 0x6E
Now, every incoming byte will be written to the internal EEPROM starting with address 0x204. Each byte will be echoed if the write was completed.
Important: Saving one byte takes about 4ms, so the next byte must be send AFTER the echo of the current byte was received from the software.
A total of 251Byte can be send. After a timeout of about 30ms, the current byte transmission will be aborted and 0xFF is send instead of the received value